## Slide 1: Other Radiographic Equipment
Amy Crane, DVM

## Slide 2: In This Lecture
Learn about x-ray beam limiting devices
Learn about grids
Know what a cassette is
Know what intensifying screens are
Know about film color sensitivity and why it is important

## Slide 3: X-ray beam limiting devices
Previously we stated that it was important to limit the size of the x-ray field
There are two reasons for this
-

## Slide 4: X-ray beam limiting devices
There are many devices used to accomplish this task of decreasing the size of the x-ray field
Cones, cylinders, and aperture diaphragms have been used in the past
Currently, the standard technology is the collimator

## Slide 5: X-ray beam limiting devices
Collimators allow a variety of x-ray field sizes
Most systems use a mirror to align the light field
Some portable systems use a laser to center the field

## Slide 6: X-ray beam limiting devices
Collimators do not focus the x-ray beam, they merely shape it by cutting off the edges of the beam.
They simply exclude the part of the beam we don’t want to use
Since x-rays do not have a mass or charge, they cannot be focused

## Slide 7: X-ray beam limiting devices
Collimators are not fool proof
Mirrors can go out of whack
A good maintenance schedule will evaluate the accuracy of the collimator light on a routine basis
Quality control test???

## Slide 8: Grids and scatter
Good collimation will reduce scatter
Tissue is, however, still exposed even though we limit the field – and scatter is produced.
Anytime scatter hits the film, there is image degradation
A grid is used to reduce scatter from hitting the film

## Slide 9: Grids
Grids should be used anytime a body part thicker than 10cm is radiographed

## Slide 10: Grids
Grid Composition
-Made of hundreds of alternating thin lead strips
with aluminum or fiber interspacers

## Slide 11: Grids
Grid Basics
-In most grids, the lead strips are angled to be in alignment with the primary x-ray beam
-This type of grid is a focused grid
-Remember, the primary beam is the part of the beam that we want to expose the film
-Scatter, however, will not be aligned with the spaces in the grid and will be absorbed by the lead strips
-Thus, scatter photons cannot hit the film because the grid absorbs them

## Slide 12: Grids
Grids are not perfect
Some of the primary beam will hit the lead strips and be absorbed
Therefore we have to account for this reduction in intensity of the primary x-ray beam when we select an exposure for a radiograph
Not all grids are the same, so be sure you know the grid factor of your particular machine.

## Slide 13: Grids
Grids are described in terms:
-Lines per inch:
-Grid Ratio:

## Slide 14: Grid Types
Parallel grids
Focused grids
Cross-hatched grids

## Slide 15: Grids
Grids are described in terms:
-Bucky factor:
-Focal range: Most grids are focused. This means that the lead strips are angled so that they are in alignment with the fan shape of the x-rays coming from the x-ray tube. These grids must be used at a precise distance from the x-ray tube to work properly. If they are used outside (too far away or too close) of the focal distance the dreaded grid cutoff occurs

## Slide 16: Grids
Grid Cutoff:

## Slide 17: Grids
Grid cutoff:
-This results in greater than normal absorption of the primary x-ray beam
-The resulting radiograph will be very light or underexposed
-The worst type of cutoff is when the grid is placed upside down. The tube side is usually labeled on the grid.

## Slide 18: Grids
Grid Lines and Moving Grids
-When grids are used properly you may see subtle images of the alternating lead strips and spacers in the radiograph
-These lines are called grid lines
-They are the unfortunate consequence of using a grid
-They are very fine and not usually distracting

## Slide 19: Grids
Fortunately we have a way of making the lines disappear by moving the grid back and forth during exposure of the radiograph
This is called a moving grid system or Potter Bucky Diaphragm

## Slide 20: Grids
Grid care:
-Grids are fragile
-If the lead strips get bent or warped it will produce artifacts that will degrade image quality
-Store flat and handle gently
-Instruct your equine patients that they should
not kick or stand on the grids!

## Slide 21: Cassettes
Cassettes:
-Rigid, light tight containers that hold the x-ray film and screens
-The screens are permanently mounted in the cassettes
-There are many sizes of cassettes available to match the commonly available sizes of film
-Many cassettes have a rectangular lead block that covers one portion of the screen. This allows for imprint identification later prior to processing

## Slide 22: Cassettes
Cassettes serve two main purposes

## Slide 23: Cassettes
You need to take care of your cassettes
Using a damaged cassette will result in a crummy radiograph due to:
-Light leakage around the edges
-Warping will hamper screen function because the screen will not be pressed tightly against the film

## Slide 24: Screens
Intensifying screens
-In general, the sensitivity of x-ray film to x-rays is poor
-X-rays can be used to expose film but it would take a lot of radiation and a very long time for most purposes
-Intensifying screens are used to increase the efficiency of film exposure

## Slide 25: Screens
Intensifying Screens:
-This increases efficiency because radiographic film is more sensitive to light than x-rays
-They are called intensifying screens because they intensify the effect that x-rays have on radiographic film
-The color of the light emitted is either blue or green light

## Slide 26: Screens
Composition
  Supporting base with a layer of fluorescent crystals
  Crystal layer is coated with a protective layer
Crystals emit visible light
  Blue or green
  When excited by x-rays
  Small flashlights
Screens are inside the cassette and in contact with the film

## Slide 27: Screens
Types of crystals
  Calcium tungstate
  Rare earth series

## Slide 28: Screens
Screen speed
  Consequence of screen usage is amount of x-ray radiation needed to cause exposure is reduced
  Film isn’t as sensitive to x-rays so it would take more radiation without screens
Different speeds of screen
  Slower speed uses higher mAs setting
  Faster speed uses lower mAs setting
What determines screen speed
  Size of crystals
  Thickness of crystal layer

## Slide 29: Screens
Basically, the more you have of a given phosphor, the faster the system will be
However, using more phosphor will greatly decrease image detail

## Slide 30: Screens
Intensifying screen benefits
-Intensifying screens vastly decrease the amount of radiation our patients and personnel receive
-Another benefit of intensifying screens is that they enable us to use faster exposures
-Intensifying screens also increase contrast

## Slide 31: Screens
Screens – putting theory into practice
-You need to select a good balance between the speed of the screen and the detail you need for your particular application
-Most screens are rare earth screens composed of phosphors from the rare-earth elements of the periodic table.

## Slide 32: Screens
FYI: Older screens used calcium tungstate as the phosphor. They are not as efficient as the rare earth crystals at converting x-rays to light. Therefore thick layers of phosphor and large phosphor particles were necessary (which decreases detail)

## Slide 33: Cassettes and Screens
Screen Maintenance
Keep clean because dust and film accumulate on the screen which will reduce performance
Clean with a commercially available screen cleaner or alcohol and a lint free cloth
Cassettes should never be left open when not in use
Screens wear out with age which decreases resolution and speed
A high-speed five year old screen has roughly the same efficiency as a new, slower speed screen

## Slide 34: The Screen System
Screens don’t do anything by themselves, they need film
So when you talk about screens you need to talk about film
Fortunately we have a whole lecture on film!!!
Keep in mind that one characteristic of x-ray film that must be considered when you purchase film is color sensitivity of the film

## Slide 35: The Screen System
X-ray films are only sensitive to one color of visible light
But, different screens give off different colors of visible light
Therefore, you must match the film color sensitivity to the light being emitted from the screen
Older screens emit a ‘blue violet’ light and appropriate film was ‘blue light’ sensitive
Newer rare earth screens emit ‘green’ light ad require ‘green’ sensitive film