## Slide 1: Artifacts and Errors

## Slide 2: Artifact
Any opacity that doesn’t correspond to an actual anatomic structure

## Slide 3: Film storage and handling artifacts
Pressure/crescents
Scratches
Fingerprints
Static
Material in cassette
Fog

## Slide 4: Pressure/crescents
Finger crescent
  Appear black and focal crease mark on film
  Improper handling of film
  Prevention
    Take time
    Handle film by edges only

## Slide 5: Scratches
Scratch in the emulsion
  Can appear black or white
  White
  Black

## Slide 6: Fingerprints
  Can be either black or white
    White more common
  Prevention
    Hands clean and dry
    Touch only film edges

## Slide 7: Static
Static electricity
  Prevention
    Remove film from box slowly
    Don’t drag film across surface as placing in cassette
    Moderate humidity
    Anti-static cassette cleaners

## Slide 8: Material in cassette/dirty screen
White and focal
Visible light from screen can’t reach and expose film emulsion

## Slide 9: Film fog
  Any unwanted film exposure turn part of film black
  Before or after exposure

## Slide 10: Light leak
Type of film fog
Looks black, usually focal

## Slide 11: Positioning artifacts
Magnification
Distortion
Grid cutoff
Patient rotation
Cassette upside down

## Slide 12: Magnification
Patient closer to film
Patient farther away from film

## Slide 13: Distortion
Patient is not parallel to film

## Slide 14: Grid cutoff
Purpose of a grid
If not centered the more of the primary beam is absorbed

## Slide 15: Patient rotation
Form of distortion

## Slide 16: Cassette upside down
Obvious how this happens

## Slide 17: Exposure artifacts
Motion
Double exposure
Overexposure
Underexposure

## Slide 18
Motion
  Blurry image
  Prevention
    Sedation or better restraint
    Closing mouth to prevent panting
    Faster time

## Slide 19
Double exposure
  Appears as blurred image
  2 identical images superimposed
  Pushed exposure button twice
  Forgot to change cassette between exposures
  Prevention
    Push button only once
    Change cassettes

## Slide 20: Other Types of Artifacts
Wet haircoat

## Slide 21: Other Types of Artifacts
Folded film in cassette

## Slide 22: Other Types of Artifacts
Contrast media on table, patient, cassette etc.

## Slide 23: Other Types of Artifacts
Portion of Film Unexposed
  Causes
    Collimation
    Lead shield
    Central x-ray beam not perpendicular to cassette
    Cassette not directly under beam (buckey tray not pushed in)

## Slide 24: Other Types of Artifacts
White Dots

## Slide 25: Quality Control

## Slide 26: Quality Control Program
Preventative maintenance
Equipment calibration
In-service education
Evaluation of new products

## Slide 27: Purpose
Provides way to minimize dose of radiation
Creates the production of quality radiographs
Its use leads to a decreased number of repeated films

## Slide 28: QC in Veterinary Radiology
Want to minimize exposure of assistants, not as concerned about the patients
Vet Tech major part of QC program
  Perform tests
  Interpret the data
  Keep records
The data collected is objective, not based on opinion or personal preference for image contrast and quality

## Slide 29: Equipment
Folder for records
Measuring tape
Thermometer
Nine pennies
Carpenters level
Screen-film contact mesh

## Slide 30: SID-source image distance
Measure from focal spot on tube housing to tabletop.
Measure from table top to top of cassette in the bucky tray
Add measurements together
Sometimes referred to as FFD – film focal distance
Notes: Source image distance
If no mark on the tube housing, divide the tube housing end cap into fourths. Use the bottom fourth as the focal spot

## Slide 31: Perpendicularity
To ensure the x-ray beam is properly centered, be sure the tube stand, collimator and x-ray tube are perpendicular and properly aligned
Use the level to confirm that the tube is level and parallel to the table
Stand at the end of the table and look at the tube, collimator and tube stand. Verify they are perpendicular
Stand alongside the table and verify the same info.
Adjust as needed

## Slide 32: X-ray Apparatus
Tube/Table Locks
X-ray field light

## Slide 33: X-ray Apparatus Cont.
To ensure the light field determined by the collimator dials is accurate

## Slide 34: X-ray Apparatus Cont.
Angulation indictor
  Place carpenters level on the tabletop (it should be level
  Place the carpenters level on the bottom of the collimator (also should be level)
  Rotate x-ray tube 15 degrees
  Using protractor measure degree of angulation *should also be 15 degrees
  Repeat at 30 and 45 degrees

## Slide 35: Viewbox Uniformity
Clean view box
If super dirty the spots could mislead interpretation

## Slide 36: Light Field/X-ray Field Alignment
Center x-ray tube over table
Set SID to 40 inches
Put cassette in bucky tray
Set collimator field indicators at a field that is approximately 6 x 8 inches
Turn collimator light on
9 penny collimation test

## Slide 37: Screen-Film Contact

## Slide 38: Calibration
Should be at least annually
  kVp
  mA
  Timer
  Filtration

## Slide 39: Darkroom QC
Needs to be clean
Check for light leaks
Check safe light for distance and bulb wattage and proper filter
If doing manual processing temperature checks